|  | AUSTRALIAN ASTRONOMICAL OBSERVATORY |
| --- | --- |

# 2dFsim

| Document Number | AAT-EL-2DFSIM-001-HI-TR |
| --- | --- |
| Version Number | 1.6 |
| Status | Draft |
| Issue Date | April, 2017 |
| Author | Vijay Nichani, Lew Waller, Chris McCowage |
| Filename | 2dFsimHardwareInterfaceTechRef V1.7 tjf_comments.docx |

# Hardware Interface

# Technical Reference

Change Record

| Version Number | Date | Section/Page | Reason/Remarks | Initials |
| --- | --- | --- | --- | --- |
| 1.0 | April 2017 |  | Original version using 2DFSIM HI TR as template | cmc |
| 1.1 | May 2017 |  | Copley details | cmc |
| 1.2 |  |  |  |  |
| 1.3 |  |  |  |  |
| 1.4 |  |  |  |  |
| 1.5 |  |  |  |  |
| 1.6 | Aug 2017 |  | Motor and encoder details | cmc |
| V1.7 |  |  |  |  |
| 1.8 |  |  |  |  |
| 1.9 |  |  |  |  |
| 1.10 |  |  |  |  |

Contents

1	Introduction	6

2	Description	6

3	IP Addresses	6

4	CAN Bus Interfaces	7

4.1	Ethernet-CAN Adapters	7

4.2	CAN Bus Devices	7

4.2.1	CAN Bus Baudrate	7

4.2.2	CANopen I/O Nodes	7

4.2.3	CANopen Servo Drive Nodes	8

4.2.4	2dFsim CAN Nodes	8

4.2.5	2dFsim Ancillary CAN Nodes	9

5	CANopen Digital and Analog I/O	10

5.1	2dFsim Control I/O	10

5.2	Calibration Lamp Control CANopen I/O (TBC)	21

6	CANopen Servo Drive	25

6.1	Drive Configuration	25

6.1.1	Servo Drive Configuration	29

6.1.1.1	2DFSIM Gantry Mechanisms	29

6.1.1.1.1	Amplifier X1, X2 & Y Motor Configuration Data	29

6.1.1.1.2	Amplifier Digital Input Configuration Data	31

6.1.1.1.3	Amplifier Digital Output Configuration Data	32

6.1.1.2	2DFSIM Gantry Mechanisms	32

6.1.1.2.1	Amplifier Theta Motor Configuration Data	32

6.1.1.2.2	Amplifier Digital Input Configuration Data	34

6.1.1.2.3	Amplifier Digital Output Configuration Data	35

6.1.1.3	2DFSIM Gantry Mechanisms	36

6.1.1.3.1	Amplifier Z Motor Configuration Data	36

6.1.1.3.2	Amplifier Digital Input Configuration Data	38

6.1.1.3.3	Amplifier Digital Output Configuration Data	39

6.1.1.4	2DFSIM Gantry Mechanisms	40

6.1.1.4.1	Amplifier Jaw Motor Configuration Data	40

6.1.1.4.2	Amplifier Digital Input Configuration Data	41

6.1.1.4.3	Amplifier Digital Output Configuration Data	42

6.1.2	EL1216 Breakout Board Configuration	44

7	Operational Descriptions	45

8	Reference Documents	45

Tables

Table 1: 2dFsim Robotics (2dFsim rack) CAN Node Addresses and Functions	8

Table 2: 2dFsim Ancillary CAN Node Addresses	9

Table 3: 2dFsim Control I/O	10

Table 4: Calibration Lamp Control CANopen I/O	21

Table 5: 2dFsim X1, X2 Motor/Encoder Part Numbers	25

Table 6: 2dFsim Y Motor/Encoder Part Numbers	26

Table 7: 2dFsim Z axis Positioner Maxon Motor/Encoder Part Numbers	26

Table 8: 2dFsim Theta Motor/Encoder Part Numbers	27

Table 9: Spectrograph Focus Actuator Maxon Motor Part Numbers	27

Table 10: Accelnet AEROTECH Motor Configuration Data	29

Table 11: Accelnet Digital Input Configuration Data (Motors X1, X2 & Y)	31

Table 13: Accelnet Plus Digital Output Configuration Data (Motors X1, X2 & Y)	32

Table 13: Accelnet UTH Motor Configuration Data	33

Table 14: Accelnet Digital Input Configuration Data (Motor Theta)	34

Table 15: Accelnet Plus Digital Output Configuration Data (Motor Theta)	35

Table 16: Accelnet MAXON Motor Configuration Data	36

Table 17: Accelnet Digital Input Configuration Data (Motor Z)	38

Table 18: Accelnet Digital Output Configuration Data (Motor Z)	39

Table 14: Accelnet Plus Maxon Motor 110164 Motor Configuration Data	40

Table 15: Accelnet Plus Digital Input Configuration Data	42

Table 16: Accelnet Plus Digital Output Configuration Data	43

Table 17: EL1216 Amplifier Breakout Board Configuration	44

# Introduction

This document describes the electronics interface to the 2dF simulator instrument control system. The 2dF simulator is an off telescope single gantry robot intended as a test bed for 2dF development and maintenance purposes. The robot consists of a complex gripper unit mounted on an X-Y Cartesian gantry, mounted above the fixed field plate mounted on a purpose built frame which can be rotated about the Y axis to produce changing a gravity vector in one axis. The gripper gantry has two X axes and a Y axis cross-beam in an 'H' configuration. The axes are driven by linear AC servo motors. Mounted on the Y axis is the gripper assembly comprising Z, theta and jaw axes. The Z and Jaw axes are driven by a brushed DC rotary motors with brake fitted to Z. The theta axis is driven by a rotary AC servo motor and the jaw by a brushed DC rotary motor.

# Description

The 2dF simulator comprises linear motors on the X1, X2, Y & Theta (T) axes and DC motors on the Z and Jaw (J). The axes are encoded by incremental encoders that require homing at startup. The X1, X2, Y & Z axis produce sin/cos waveforms while Theta and Jaw produce two digital waveforms in quadrature. Homing marks are present on X1, X2, Y and T whilst the Z and J axes home off of their positive limit sensors, upward and opening directions respectively. Brakes are intended to be fitted to X1, X2 & Y. Z has a brake fitted to the drive motor which has to be powered to release. Theta is connected via slip rings and capable of continuous rotation. Drive is provided by Copley Accelnet amplifiers and associated power supply. The Copley amplifiers are on a network implemented using canbus. Communication with the control computer is via Analytica AnaGate CAN uno Din rail Ethernet to CANopen adapters. The control software will consist of components of the 2df control system via a new hardware abstraction layer to communicate with the new amplifiers and associated I/O systems.

# IP Addresses

2dFsim uses equipment with network interfaces that require IP addresses.

All network equipment automatically acquires their address using DHCP when they are plugged in and powered on at the AAT.

The AAT IP addresses for 2dFsim network devices may be found in [RD3].

# CAN Bus Interfaces

## Ethernet-CAN Adapters

The 2DFSIM instrument has several CAN buses. These are implemented with Ethernet-CAN adapters (Analytica AnaGate CAN uno Din rail).

The following adapters are used:

- The 2dFsim Ethernet-CAN adapter controls CAN bus devices in the 2dFsim instrument control cabinet. These are described in §4.2.5.

- It is undecided if there will be an I/O Control Ethernet-CAN adapter to control a CANopen I/O module in the 2dFsim instrument control cabinet. However it is currently intended that there will be one in deployed in the upgraded operational 2dF robot and it would be prudent to allow for multiple I/O Control Ethernet-CAN adapters in the architecture to control I/O, calibration lamps, ADC etc. This is described in 5.1

High level software controls all CAN bus devices using Copley Motion Libraries (CML).

## CAN Bus Devices

The following sections describe the devices on each of the CAN buses.

### CAN Bus Baudrate

All CAN bus servo drives have their CAN bus baudrate set to 500,000 bits per second (§4.2.3). CAN bus I/O nodes set their baudrate automatically after power on (§4.2.2).

### CANopen I/O Nodes

CANopen I/O is implemented using Beckhoff Bus Terminal modules and CANopen bus couplers.

The BK5150 CANopen bus coupler has its node address set by rotary switches on the unit.

The bus coupler determines the CAN bus baud rate using an autobaud technique that monitors bus traffic after the unit is powered on.

### CANopen Servo Drive Nodes

Copley Controls Accelnet Plus 1-Axis Panel (ADP-180-18-S) servo drives are used to control X1, X2 and Y servo motors.

Copley Controls Accelnet Plus 1-Axis Panel (ADP-180-09) servo drives are used to control Z, Theta and Jaw servo motors.

The Accelnet Plus Panel has its CAN bus node address set by rotary switches on the unit.

The Accelnet Plus Panel’s CAN bus baud rate is programmed in Flash memory using the CME application.

### 2dFsim CAN Nodes

The following table describes the CAN bus node addresses of the 2dFsim CAN bus:

Table 1: 2dFsim Robotics (2dFsim rack) CAN Node Addresses and Functions

| Address |  | Device | Function | Motor Id |
| --- | --- | --- | --- | --- |
| decimal | hexadecimal |  |  |  |
| 1 | 0x01 | CANopen I/O | 2dFsim Gantry Control I/O |  |
| 11 | 0x0b | CANopen servo drive | 2dFsim Gantry X1 (Motor 1) | M1 |
| 12 | 0x0c | CANopen servo drive | 2dFsim Gantry X1 (Motor 2) | M2 |
| 13 | 0x0d | CANopen servo drive | 2dFsim Gantry Y (Motor 3) | M3 |
| 14 | 0x0e | CANopen servo drive | 2dFsim Gantry Z | M4 |
| 15 | 0x0f | CANopen servo drive | 2dFsim Gantry Theta | M5 |
| 16 | 0x10 | CANopen servo drive | 2dFsim Gantry Jaw | M6 |

Descriptions of the operations of these actuators are given in [RD1].

Note that motors currently motor 2 is not directly controlled. M2 is configured to operate as a “slave” to motors 1, which operate as the “master” with M1 encoder output as the demanded position on M2. n.b. It has not been determined if this configuration will be suitable in the application and the behavior and performance will need to be determined operating under the CML. Please allow both configurations in the architecture with a mechanism to easily switch between the two modes:

- M1 (X1) as the X axis with M2 (X2) “slaved” to M1 (X1)

- M1 (X1) and M2 (X2) as independent axes “bound” into the X co-ordinate system

### 2dFsim Ancillary CAN Nodes

The following table describes the CAN bus node addresses of the Instrument Rack CAN bus:

Table 2: 2dFsim Ancillary CAN Node Addresses

| Address |  | Device | Function |
| --- | --- | --- | --- |
| decimal | hexadecimal |  |  |
| 1 | 0x01 | CANopen I/O | Power Control I/O, Back illumination and Plate Fiducial Control. |

It has not been decided if there is a requirement for remote control of the power supplies, back illumination and plate fiducial control on the 2dFsim. Please provide configurable option to allow control of the 2dFsim power supplies via a digital I/O or manual power supply control.

# CANopen Digital and Analog I/O

The following tables describe inputs and outputs connected to the CANopen I/O modules used in 2dFsim. These are used for controlling the 2dFsim ancillary hardware, the system power supplies, back illumination and the plate fiducials.

## 2dFsim Control I/O

This I/O module is located in the 2dFsim Control rack on the 2dFsim trolley.

Table 3: 2dFsim Control I/O

| Node Name | CAN Node Address | Group or Channel | Bit in Group | Input or output | Type | Function | Description |  | CANopen Access Method |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | Inactive State | Active State |  |
| 2dFsim CANopen I/O | 1 | 0 | 0 | 1 | Digital Input | Instrument Enabled(/Disabled) Status | The instrument is enabled. | The instrument is disabled. | TxPDO1 |
|  |  |  | 1 | 2 | Digital Input | Emergency Stop Status | The instrument enable system is in the emergency stopped state (an emergency stop switch is active or has been released). | The instrument enable system is not in the emergency stopped state. |  |
|  |  |  | 2 | 3 | Digital Input | Manual Enable Switch | Manual Switch Input is in Enabled position | Manual Switch Input is in Disabled position |  |
|  |  |  | 3 | 4 |  | Over Ride Switch Input | Override Switch is Off | Override Switch is ON. This allows operating in the override position. Note this will also indicate with flashing Orange Light. Status Override LED in the front panel will also turn ON. This mode is allowed during maintenance. |  |
|  |  |  | 4 | 5 | Digital Input | Dry air pressure | Indicates dry air pressure OK | Dry air pressure low |  |
|  |  |  | 5 | 6 | Digital Input | Cover Switch | Cover is Open. This will Trigger an E-Stop | Cover is Closed |  |
|  |  |  | 6 | 7 | Digital Input | Cover Over Ride Switch | Indicates Cover Over Ride Switch is ON | Indicates Cover Over Ride Switch is OFF |  |
|  |  |  | 7 | 8 | Digital Input | Un Allocated |  |  |  |
|  |  | 0 | 0 | 1 | Digital Output | Power to 2dFsim servo power supply | 2dFsim servo power supply OFF | 2dFsim servo power supply ON | RxPDO1 |
|  |  |  | 1 | 2 | Digital Output | Back illumination control | Back illumination OFF | Back illumination ON. |  |
|  |  |  | 2 | 3 | Digital Output | Plate Fiducials | Plate fiducials are turned OFF | Plate fiducials are turned ON. |  |
|  |  |  | 3 | 4 | Digital Output | Power to Gripper Camera | Gripper Camera Power is Turned OFF | Gripper Camera Power is Turned ON |  |
|  |  |  | 4 | 5 | Digital Output | Hazard Warning Lights | Hazard Warning lights OFF | Hazard Warning lights ON |  |
|  |  |  | 5 | 6 | Digital Output | Audible Warning | Audible Warning OFF | Audible Warning ON |  |
|  |  |  | 6 | 7 | Digital Output | Cover Over Ride Light | Indicates Cover Over Ride Light is ON | Indicates Cover Over Ride Light is OFF |  |
|  |  |  | 7 | 8 | Digital Output | Dry air pressure | Indicates dry air pressure OK | Dry air pressure low |  |
|  |  | 1 | 0 | 1 | Digital Output | Software Emergency Stop | The software emergency stop relay is de-energised. | Energise the software emergency stop relay. Note that the relay must be de-energised before the instrument enable system can be put back in the enabled state. Manual intervention is required to re-enable the instrument. |  |
|  |  |  | 1 | 2 | Digital Output | Software Status Run-Mode LED | Indicates Software is not in Run Mode operation. Status RUN Mode LED is OFF | Indicates Software is in Run Mode operation. Status RUN Mode LED is ON. |  |
|  |  |  | 2 | 3 | Digital Output | Software Status System Error LED | Indicates Software has not encountered any Error. Error LED is OFF. | Indicates Software has encountered an Error. Error LED is ON. |  |
|  |  |  | 3 | 4 | Digital Output | Software is running in Over Ride mode | Override Status LED is OFF. | Override Status LED is ON. Indicating system has detected an error such Cover Open, but an Over Ride switch has been enabled and software is running in override mode. |  |
|  |  |  | 4 | 5 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 5 | 6 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 6 | 7 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 7 | 8 | Digital Output | Un Allocated |  |  |  |
|  |  | 2 | 0 | 1 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 1 | 2 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 2 | 3 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 3 | 4 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 4 | 5 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 5 | 6 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 6 | 7 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 7 | 8 | Digital Output | Un Allocated |  |  |  |
|  |  | 3 | 0 | 1 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 1 | 2 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 2 | 3 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 3 | 4 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 4 | 5 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 5 | 6 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 6 | 7 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 7 | 8 | Digital Output | Un Allocated |  |  |  |
|  |  | 4 | 0 | 1 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 1 | 2 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 2 | 3 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 3 | 4 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 4 | 5 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 5 | 6 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 6 | 7 | Digital Output | Un Allocated |  |  |  |
|  |  |  | 7 | 8 | Digital Output | Un Allocated |  |  |  |
|  |  | 0 | 16-Bit | 1 | Thermocouple (Analog) Input | Temperature Sensor1 | 16-Bit Value which may need calibration but Typical Value of Tc= 22°C get 0x00E3 and Tc=31°C get 0x0164. Value is 2’s complement in range of -32768 to 32767. Resolution is 0.1 °C per digit. Location Field Plate temperature. |  | TxPDO2 |
|  |  | 1 | 16-Bit | 2 | Thermocouple (Analog) Input | Temperature Sensor2 | 16-Bit Value which may need calibration but Typical Value of Tc= 22°C get 0x00E3 and Tc=31°C get 0x0164. Value is 2’s complement in range of -32768 to 32767. Resolution is 0.1 °C per digit. Location Gantry Steel Frame Temperature. |  |  |
|  |  | 2 | 16-Bit | 3 | Thermocouple (Analog) Input | Temperature Sensor3 | 16-Bit Value which may need calibration but Typical Value of Tc= 22°C get 0x00E3 and Tc=31°C get 0x0164. Value is 2’s complement in range of -32768 to 32767. Resolution is 0.1 °C per digit. Location Gantry Y Beam Box Section Temperature. |  |  |
|  |  | 3 | 16-Bit | 4 | Thermocouple (Analog) Input | Temperature Sensor4 | 16-Bit Value which may need calibration but Typical Value of Tc= 22°C get 0x00E3 and Tc=31°C get 0x0164. Value is 2’s complement in range of -32768 to 32767. Resolution is 0.1 °C per digit. Location 2dFSim electronics rack able to a general purpose monitoring channel and for diagnostic monitoring of points including amplifier heat sinks. Reading here is indication on the temperature of the various amplifiers and power supplies. Warning needs to be generated if Temperature goes above 45°C & Error if it goes above 48°C. |  |  |
|  |  | 4 | 16-Bit | 1 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  | TxPDO3 |
|  |  | 5 | 16-Bit | 2 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |
|  |  | 6 | 16-Bit | 3 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |
|  |  | 7 | 16-Bit | 4 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |
|  |  | 8 | 16-Bit | 1 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  | TxPDO4 |
|  |  | 9 | 16-Bit | 2 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |
|  |  | 10 | 16-Bit | 3 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |
|  |  | 11 | 16-Bit | 4 | Analog Input – 4..20 mA | Un Allocated | Not connected – reads 0. |  |  |

## Calibration Lamp Control CANopen I/O (TBC)

The Calibration Lamp Control is not required for the 2dFsim however it may be prudent to make provision for a CANopen I/O module Calibration Lamp Control for when the control system migrates to the observing top end. The 2dF calibration system is currently located in the sky baffle with two flaps to provide a light seal and a reflective screen. However there is a proposed general purpose calibration system to be mounted on the telescope top ring to illuminate the windscreen mounted calibration reflective screen. (TBC)

Table 4: Calibration Lamp Control CANopen I/O

| Node Name | CAN Node Address | Group or Channel | Bit in Group | Input or output | Type | Function | Description |  | CANopen Access Method |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | Inactive State | Active State |  |
| Calibration Lamp Control CANopen I/O | 1 | 0 | 0 | 1 | Digital Input | Lamps Enabled(/Disabled) Status | The calibration lamp unit is disabled. | The calibration lamp unit is enabled. | TxPDO1 |
|  |  |  | 1 | 2 | Digital Input | Unallocated |  |  |  |
|  |  |  | 2 | 3 | Digital Input | Unallocated |  |  |  |
|  |  |  | 3 | 4 | Digital Input | Unallocated |  |  |  |
|  |  |  | 4 | 5 | Digital Input | Unallocated |  |  |  |
|  |  |  | 5 | 6 | Digital Input | Unallocated |  |  |  |
|  |  |  | 6 | 7 | Digital Input | Unallocated |  |  |  |
|  |  |  | 7 | 8 | Digital Input | Unallocated |  |  |  |
|  |  | 0 | 0 | 1 | Digital Output | CuAr Lamp 1 On/Off | Lamp Off | Lamp On | RxPDO1 |
|  |  |  | 1 | 2 | Digital Output | CuAr Lamp 2 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 2 | 3 | Digital Output | FeAr Lamp 1 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 3 | 4 | Digital Output | FeAr Lamp 2 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 4 | 5 | Digital Output | ThAr Lamp 1 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 5 | 6 | Digital Output | ThAr Lamp 2 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 6 | 7 | Digital Output | He Lamp 1 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 7 | 8 | Digital Output | Ne Lamp 2 On/Off | Lamp Off | Lamp On |  |
|  |  | 1 | 0 | 1 | Digital Output | QTH Lamp 1 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 1 | 2 | Digital Output | QTH Lamp 2 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 2 | 3 | Digital Output | QTH Lamp 3 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 3 | 4 | Digital Output | QTH Lamp 4 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 4 | 5 | Digital Output | QTH Lamp 5 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 5 | 6 | Digital Output | QTH Lamp 6 On/Off | Lamp Off | Lamp On |  |
|  |  |  | 6 | 7 | Digital Output | Unallocated |  |  |  |
|  |  |  | 7 | 8 | Digital Output | Unallocated |  |  |  |

# CANopen Servo Drive

## Drive Configuration

This section contains configuration information for the Copley Controls Accelnet ADP-180-18-S (sin/cos) and ADP-180-09 (Quadrature) CANopen servo drives used in 2DFSIM.

The gantry X1, X2 are of the same type and consist of the following components:

Table 5: 2dFsim X1, X2 Motor/Encoder Part Numbers

| AEROTECH Motor Part No. | Component | Name | Comment |
| --- | --- | --- | --- |
| BLM-386-A-AC-H-S-5000 | Motor | U-channel BLM series non-magnetic forcer |  |
| MT600P | Magnet Track | Magnet Track | Length 24” Magnetic Pole Pitch 1.2” |
| MT480P | Magnet Track | Magnet Track | Length 19.2” Magnetic Pole Pitch 1.2” |
| LIS73-1 B121 F00720 COKA | Encoder | NUMERIC JENA LIS 71-1 | This encoder is made by Numeric Jena and is a linear incremental sealed 720 mm DOUBLEFLEX steel tape with a Grating Pitch of 20 µM 1 Vpp Sinewave 5 µM Resolution Accuracy class ±2 µM Reference marks Distance coded at 1000 x GP |
| 301212 | TBD |  |  |

Table 6: 2dFsim Y Motor/Encoder Part Numbers

| AEROTECH Motor Part No. | Component | Name | Comment |
| --- | --- | --- | --- |
| BLM-386-A-AC-H-S-5000 | Motor | U-channel BLM series non-magnetic forcer |  |
| LIS73-1 B121 F00720 COKA | Encoder | NUMERIC JENA LIS 71-1 | This encoder is made by Numeric Jena and is a linear incremental sealed 720 mm DOUBLEFLEX steel tape with a Grating Pitch of 20 µM 1 Vpp Sinewave 5 µM Resolution Accuracy class ±2 µM Reference marks Distance coded at 1000 x GP |
| MT600P | Magnet Track | Magnet Track | Length 24” Magnetic Pole Pitch 1.2” |
| 301212 | TBD |  |  |

Table 7: 2dFsim Z axis Positioner Maxon Motor/Encoder Part Numbers

| Maxon Motor Part No. | Component | Name | Comment |
| --- | --- | --- | --- |
| 273754 | Motor | RE35, 35mm Dia, Graphite Brushes, 90W, 42V |  |
| 351942 | Gearhead | GP32BZ Planetary Gearhead, Low Backlash | 5.2:1 |
| 228384 | Brake | AB28 Brake, 24Vdc, 0.4Nm |  |
| LIA20-L301-BZ | Encoder | Encoder Linear incremental Grating Period 20 µM RS422 square wave with interpolation 5x | Scale 90 mm Steel Tape SINGLEFLEX ±3 µM Reference Mark in centre of scale. |

Table 8: 2dFsim Theta Motor/Encoder Part Numbers

| UTH Motor Part No. | Component | Name | Comment |
| --- | --- | --- | --- |
| UTH-92-A-28-A-H-000 | Motor | Brushless High Torque Motor Kit | Rotor & Stator assemblies with internal halls |
| AVAGO HEDS-9140#A00 | Encoder | AVAGO HEDS-9140#A00 | This encoder is made by Avago Technologies. HEDS-9140#A00 (notionally 500 CPR) 3-channel encoder module and a custom printed encoder wheel (encnegRS274xV5.gbr) |

The 2DFSIM spectrograph actuators consist of the following components, for focus actuation and back illuminator positioning:

Table 9: Spectrograph Focus Actuator Maxon Motor Part Numbers

| Maxon Motor Part No. | Component | Name | Comment |
| --- | --- | --- | --- |
| 110160 | Motor | A-max 22 Ø22 mm, brushed, 6 Watt | 12V DC |
| AVAGO HEDS-9140#A00 | Encoder | AVAGO HEDS-9140#A00 | This encoder is made by Avago Technologies. HEDS-9140#A00 (500 CPR) 3-channel encoder module and a HEDS5140-A11: Code Wheel, 3 channel. |
| 110355 | Gearhead | Planetary Gearhead GP 22 B ∅22 mm, 0.1–0.3 Nm | 4.4 : 1 (57⁄13) |

### Servo Drive Configuration

Servo drive configuration is done using Copley Motion Explorer (CME). These tables are not complete as yet for X1, X2, Y, Theta and Z. Please refer to the Reference Documents 3,4,5 & 6 in the meanwhile.

#### 2DFSIM Gantry Mechanisms

##### Amplifier X1, X2 & Y Motor Configuration Data

The following table shows the motor configuration data for the AEROTECH motors used in the 2DFSIM Gantry as required by the Accelnet.

Table 10: Accelnet AEROTECH Motor Configuration Data

| Parameter | Motor Data |  | Copley Controls Data |  | Description |
| --- | --- | --- | --- | --- | --- |
| Manufacturer |  |  |  |  | Motor manufacturer’s name. Saved for reference in the motor data file. |
| Model No. |  |  |  |  | Motor model number. Saved for reference in the motor data file. |
| Units |  |  | Metric |  | Selects whether the parameters entered in this screen are in Metric or English units. |
| Motor Inertia |  | g.cm2 |  | kg.cm2 | The rotor inertia of the motor. Used for calculating initial velocity loop tuning values. Range: 0.00001 to 4,294 kg cm2. Default: 0.00001 kg-cm2. |
| Number of Pole Pairs |  |  |  |  |  |
| Number of Poles |  |  |  |  | (Brushless only.) The number of magnetic poles in the motor. Required for correct commutation of the motor. If the number of poles is not known, Verify the motor’s pole count using the Data accuracy test.. Range: 2 to 200. Default: 4. |
| Peak Torque |  | mN.m |  | N.m | The peak torque that the motor can produce. Peak Torque divided by torque constant = motor’s peak current limit. Range: 0.0001 to 2,100 Nm. Default: 0.0001 Nm. |
| Continuous Torque |  | mN.m |  | N.m | The continuous torque that the motor can produce. Used with the torque constant to calculate continuous current. Range: 0.0001 to 1,000 Nm. Default: 0.0001 Nm. |
| Velocity Limit |  | Rpm |  | rpm | Maximum speed of the motor. Used to calculate the velocity and acceleration limits for the velocity loop. Range dependent on encoder resolution. |
| Torque Constant |  | mN.m/A |  | N.m/Apk | Relates the motor's input current to torque produced. Sometimes abbreviated as Kt. Range: 0.001 to 1,000 Nm/Apk. Default: 0.001 Nm/Apk. |
| Speed Constant |  | rpm/V |  | krpm/V |  |
| Back emf Constant |  |  |  | V/krpm | Relates the motor's input voltage to speed. Sometimes abbreviated as Ke. Used for calculating the maximum velocity for a given amplifier bus voltage. Range: 0.01 to 21,000,000 V/Krpm. Default: 0.01 V/Krpm. |
| Resistance |  | Ω |  | Ω | Motor resistance line-to-line. Used for calculating the initial current loop tuning values. Range: 0.01 to 327Ω. Default: 0.01Ω. |
| Inductance |  | mH |  | mH | Motor inductance line-to-line. Used for calculating the initial current loop tuning values. Range: see the hardware documentation. |

##### Amplifier Digital Input Configuration Data

The following table shows the digital input configuration data for the amplifier in 2DFSIM.

Table 11: Accelnet Digital Input Configuration Data (Motors X1, X2 & Y)

| Amp. Input | Amp. Configuration |  |  | EL1216 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Pullup | Function | Action | Function | Designator | Other |
| IN1 | Yes | Amp Enable | LO Enables with Clear Faults | Enable Switch | SW1 | Up to enable |
| IN2 | Yes | Motion Abort | Active LO | Abort Switch | SW3 | Up to abort |
| IN3 | Yes | Not Configured |  |  |  |  |
| IN4 | Yes | POS Limit | HI Inhibits | Positive Limit Switch |  | For N/C switch |
| IN5 | Yes | Not Configured |  |  |  |  |
| IN6 | Yes | Not Configured |  |  |  |  |
| IN7 |  | Not Configured |  |  |  |  |
| IN8 |  | Not Configured |  |  |  |  |
| IN9 |  | Not Configured |  |  |  | Single Ended |
| IN10 |  | Not Configured |  |  |  | Single Ended |
| IN11 |  | Not Configured |  |  |  |  |
| IN12 |  | Not Configured |  |  |  |  |

##### Amplifier Digital Output Configuration Data

The following table shows the digital output configuration data for the amplifier when used with the AAO EL1216 servo amplifier breakout board in 2DFSIM.

Table 12: Accelnet Plus Digital Output Configuration Data (Motors X1, X2 & Y)

| Amp. Output | Amp. Function | Configuration | Level | EL1216 Function |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Function | Designator | Other |
| OUT1 |  | Program Control Active Low | Hi |  |  |  |
| OUT2 |  | Brake-Active High | Lo |  |  |  |
| OUT3 |  | Program Control Active Low | Hi |  |  |  |

#### 2DFSIM Gantry Mechanisms

##### Amplifier Theta Motor Configuration Data

The following table shows the motor configuration data for the UTH motor used in the 2DFSIM Gantry as required by the Accelnet.

Table 13: Accelnet UTH Motor Configuration Data

| Parameter | Motor Data |  | Copley Controls Data |  | Description |
| --- | --- | --- | --- | --- | --- |
| Manufacturer |  |  |  |  | Motor manufacturer’s name. Saved for reference in the motor data file. |
| Model No. |  |  |  |  | Motor model number. Saved for reference in the motor data file. |
| Units |  |  | Metric |  | Selects whether the parameters entered in this screen are in Metric or English units. |
| Motor Inertia |  | g.cm2 |  | kg.cm2 | The rotor inertia of the motor. Used for calculating initial velocity loop tuning values. Range: 0.00001 to 4,294 kg cm2. Default: 0.00001 kg-cm2. |
| Number of Pole Pairs |  |  |  |  |  |
| Number of Poles |  |  |  |  | (Brushless only.) The number of magnetic poles in the motor. Required for correct commutation of the motor. If the number of poles is not known, Verify the motor’s pole count using the Data accuracy test.. Range: 2 to 200. Default: 4. |
| Peak Torque |  | mN.m |  | N.m | The peak torque that the motor can produce. Peak Torque divided by torque constant = motor’s peak current limit. Range: 0.0001 to 2,100 Nm. Default: 0.0001 Nm. |
| Continuous Torque |  | mN.m |  | N.m | The continuous torque that the motor can produce. Used with the torque constant to calculate continuous current. Range: 0.0001 to 1,000 Nm. Default: 0.0001 Nm. |
| Velocity Limit |  | Rpm |  | rpm | Maximum speed of the motor. Used to calculate the velocity and acceleration limits for the velocity loop. Range dependent on encoder resolution. |
| Torque Constant |  | mN.m/A |  | N.m/Apk | Relates the motor's input current to torque produced. Sometimes abbreviated as Kt. Range: 0.001 to 1,000 Nm/Apk. Default: 0.001 Nm/Apk. |
| Speed Constant |  | rpm/V |  | krpm/V |  |
| Back emf Constant |  |  |  | V/krpm | Relates the motor's input voltage to speed. Sometimes abbreviated as Ke. Used for calculating the maximum velocity for a given amplifier bus voltage. Range: 0.01 to 21,000,000 V/Krpm. Default: 0.01 V/Krpm. |
| Resistance |  | Ω |  | Ω | Motor resistance line-to-line. Used for calculating the initial current loop tuning values. Range: 0.01 to 327Ω. Default: 0.01Ω. |
| Inductance |  | mH |  | mH | Motor inductance line-to-line. Used for calculating the initial current loop tuning values. Range: see the hardware documentation. |

##### Amplifier Digital Input Configuration Data

The following table shows the digital input configuration data for the amplifier in 2DFSIM.

Table 14: Accelnet Digital Input Configuration Data (Motor Theta)

| Amp. Input | Amp. Configuration |  |  | EL1216 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Pullup | Function | Action | Function | Designator | Other |
| IN1 | Yes | Amp Enable | LO Enables with Clear Faults | Enable Switch | SW1 | Up to enable |
| IN2 | Yes | Motion Abort | Active LO | Abort Switch | SW3 | Up to abort |
| IN3 | Yes | Not Configured |  |  |  |  |
| IN4 | Yes | POS Limit | HI Inhibits | Positive Limit Switch |  | For N/C switch |
| IN5 | Yes | Not Configured |  |  |  |  |
| IN6 | Yes | Not Configured |  |  |  |  |
| IN7 |  | Not Configured |  |  |  |  |
| IN8 |  | Not Configured |  |  |  |  |
| IN9 |  | Not Configured |  |  |  | Single Ended |
| IN10 |  | Not Configured |  |  |  | Single Ended |
| IN11 |  | Not Configured |  |  |  |  |
| IN12 |  | Not Configured |  |  |  |  |

##### Amplifier Digital Output Configuration Data

The following table shows the digital output configuration data for the amplifier when used in 2DFSIM.

Table 15: Accelnet Plus Digital Output Configuration Data (Motor Theta)

| Amp. Output | Amp. Function | Configuration | Level | EL1216 Function |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Function | Designator | Other |
| OUT1 |  | Program Control Active Low | Hi |  |  |  |
| OUT2 |  | Brake-Active High | Lo |  |  |  |
| OUT3 |  | Program Control Active Low | Hi |  |  |  |

#### 2DFSIM Gantry Mechanisms

##### Amplifier Z Motor Configuration Data

The following table shows the motor configuration data for the MAXON motors used in the 2DFSIM Gantry as required by the Accelnet.

Table 16: Accelnet MAXON Motor Configuration Data

| Parameter | Motor Data |  | Copley Controls Data |  | Description |
| --- | --- | --- | --- | --- | --- |
| Manufacturer |  |  |  |  | Motor manufacturer’s name. Saved for reference in the motor data file. |
| Model No. |  |  |  |  | Motor model number. Saved for reference in the motor data file. |
| Units |  |  | Metric |  | Selects whether the parameters entered in this screen are in Metric or English units. |
| Motor Inertia |  | g.cm2 |  | kg.cm2 | The rotor inertia of the motor. Used for calculating initial velocity loop tuning values. Range: 0.00001 to 4,294 kg cm2. Default: 0.00001 kg-cm2. |
| Number of Pole Pairs |  |  |  |  |  |
| Number of Poles |  |  |  |  | (Brushless only.) The number of magnetic poles in the motor. Required for correct commutation of the motor. If the number of poles is not known, Verify the motor’s pole count using the Data accuracy test.. Range: 2 to 200. Default: 4. |
| Peak Torque |  | mN.m |  | N.m | The peak torque that the motor can produce. Peak Torque divided by torque constant = motor’s peak current limit. Range: 0.0001 to 2,100 Nm. Default: 0.0001 Nm. |
| Continuous Torque |  | mN.m |  | N.m | The continuous torque that the motor can produce. Used with the torque constant to calculate continuous current. Range: 0.0001 to 1,000 Nm. Default: 0.0001 Nm. |
| Velocity Limit |  | Rpm |  | rpm | Maximum speed of the motor. Used to calculate the velocity and acceleration limits for the velocity loop. Range dependent on encoder resolution. |
| Torque Constant |  | mN.m/A |  | N.m/Apk | Relates the motor's input current to torque produced. Sometimes abbreviated as Kt. Range: 0.001 to 1,000 Nm/Apk. Default: 0.001 Nm/Apk. |
| Speed Constant |  | rpm/V |  | krpm/V |  |
| Back emf Constant |  |  |  | V/krpm | Relates the motor's input voltage to speed. Sometimes abbreviated as Ke. Used for calculating the maximum velocity for a given amplifier bus voltage. Range: 0.01 to 21,000,000 V/Krpm. Default: 0.01 V/Krpm. |
| Resistance |  | Ω |  | Ω | Motor resistance line-to-line. Used for calculating the initial current loop tuning values. Range: 0.01 to 327Ω. Default: 0.01Ω. |
| Inductance |  | mH |  | mH | Motor inductance line-to-line. Used for calculating the initial current loop tuning values. Range: see the hardware documentation. |

##### Amplifier Digital Input Configuration Data

The following table shows the digital input configuration data for the amplifier in 2DFSIM.

Table 17: Accelnet Digital Input Configuration Data (Motor Z)

| Amp. Input | Amp. Configuration |  |  | EL1216 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Pullup | Function | Action | Function | Designator | Other |
| IN1 | Yes | Amp Enable | LO Enables with Clear Faults | Enable Switch | SW1 | Up to enable |
| IN2 | Yes | Motion Abort | Active LO | Abort Switch | SW3 | Up to abort |
| IN3 | Yes | Not Configured |  |  |  |  |
| IN4 | Yes | POS Limit | HI Inhibits | Positive Limit Switch |  | For N/C switch |
| IN5 | Yes | Not Configured |  |  |  |  |
| IN6 | Yes | Not Configured |  |  |  |  |
| IN7 |  | Not Configured |  |  |  |  |
| IN8 |  | Not Configured |  |  |  |  |
| IN9 |  | Not Configured |  |  |  | Single Ended |
| IN10 |  | Not Configured |  |  |  | Single Ended |
| IN11 |  | Not Configured |  |  |  |  |
| IN12 |  | Not Configured |  |  |  |  |

##### Amplifier Digital Output Configuration Data

The following table shows the digital output configuration data for the amplifier when used n 2DFSIM.

Table 18: Accelnet Digital Output Configuration Data (Motor Z)

| Amp. Output | Amp. Function | Configuration | Level | EL1216 Function |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Function | Designator | Other |
| OUT1 |  | Program Control Active Low | Hi |  |  |  |
| OUT2 |  | Brake-Active High | Lo |  |  |  |
| OUT3 |  | Program Control Active Low | Hi |  |  |  |

#### 2DFSIM Gantry Mechanisms

##### Amplifier Jaw Motor Configuration Data

The following table shows the motor configuration data for the MAXON motors used in the 2DFSIM Gantry as required by the Accelnet.

Table 19: Accelnet Plus Maxon Motor 110164 Motor Configuration Data

| Parameter | Maxon Motor Data |  | Copley Controls Data |  | Description |
| --- | --- | --- | --- | --- | --- |
| Manufacturer | Maxon Motor |  | Maxon Motor |  | Motor manufacturer’s name. Saved for reference in the motor data file. |
| Model No. | A-max 22, 110164 |  | A-max 22, 110164 |  | Motor model number. Saved for reference in the motor data file. |
| Units |  |  | Metric |  | Selects whether the parameters entered in this screen are in Metric or English units. |
| Motor Inertia | 4.2 | g.cm2 | 0.0042 | kg.cm2 | The rotor inertia of the motor. Used for calculating initial velocity loop tuning values. Range: 0.00001 to 4,294 kg cm2. Default: 0.00001 kg-cm2. |
| Number of Pole Pairs |  |  |  |  |  |
| Number of Poles |  |  |  |  | (Brushless only.) The number of magnetic poles in the motor. Required for correct commutation of the motor. If the number of poles is not known, Verify the motor’s pole count using the Data accuracy test.. Range: 2 to 200. Default: 4. |
| Peak Torque | 23.7 | mN.m | 0.0237 | N.m | The peak torque that the motor can produce. Peak Torque divided by torque constant = motor’s peak current limit. Range: 0.0001 to 2,100 Nm. Default: 0.0001 Nm. |
| Continuous Torque | 6.9 | mN.m | 0.0069 | N.m | The continuous torque that the motor can produce. Used with the torque constant to calculate continuous current. Range: 0.0001 to 1,000 Nm. Default: 0.0001 Nm. |
| Velocity Limit | 10500 | Rpm | 10500 | rpm | Maximum speed of the motor. Used to calculate the velocity and acceleration limits for the velocity loop. Range dependent on encoder resolution. |
| Torque Constant | 21.2 | mN.m/A | 0.00212 | N.m/Apk | Relates the motor's input current to torque produced. Sometimes abbreviated as Kt. Range: 0.001 to 1,000 Nm/Apk. Default: 0.001 Nm/Apk. |
| Speed Constant | 450 | rpm/V | 0.450 | krpm/V |  |
| Back emf Constant |  |  | 2.22 | V/krpm | Relates the motor's input voltage to speed. Sometimes abbreviated as Ke. Used for calculating the maximum velocity for a given amplifier bus voltage. Range: 0.01 to 21,000,000 V/Krpm. Default: 0.01 V/Krpm. |
| Resistance | 3.44 | Ω | 21.5 | Ω | Motor resistance line-to-line. Used for calculating the initial current loop tuning values. Range: 0.01 to 327Ω. Default: 0.01Ω. |
| Inductance | 1.37 | mH | 1.37 | mH | Motor inductance line-to-line. Used for calculating the initial current loop tuning values. Range: see the hardware documentation. |

##### Amplifier Digital Input Configuration Data

The following table shows the digital input configuration data for the amplifier when used in 2DFSIM.

Table 20: Accelnet Plus Digital Input Configuration Data

| Amp. Input | Amp. Configuration |  |  | EL1216 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Pullup | Function | Action | Function | Designator | Other |
| IN1 | Yes | Amp Enable | LO Enables with Clear Faults | Enable Switch | SW1 | Up to enable |
| IN2 | Yes | Motion Abort | Active LO | Abort Switch | SW3 | Up to abort |
| IN3 | Yes | Not Configured |  |  |  |  |
| IN4 | Yes | POS Limit | HI Inhibits | Positive Limit Switch |  | For N/C switch |
| IN5 | Yes | Not Configured |  |  |  |  |
| IN6 | Yes | Not Configured |  |  |  |  |
| IN7 |  | Not Configured |  |  |  |  |
| IN8 |  | Not Configured |  |  |  |  |
| IN9 |  | Not Configured |  |  |  | Single Ended |
| IN10 |  | Not Configured |  |  |  | Single Ended |
| IN11 |  | Not Configured |  |  |  |  |
| IN12 |  | Not Configured |  |  |  |  |

##### Amplifier Digital Output Configuration Data

The following table shows the digital output configuration data for the amplifier when used in 2DFSIM.

Table 21: Accelnet Plus Digital Output Configuration Data

| Amp. Output | Amp. Function | Configuration | Level | EL1216 Function |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Function | Designator | Other |
| OUT1 |  | Program Control Active Low | Hi |  |  |  |
| OUT2 |  | Brake-Active High | Lo |  |  |  |
| OUT3 |  | Program Control Active Low | Hi |  |  |  |

### EL1216 Breakout Board Configuration

It is yet to be decided if the EL1216 Breakout Board Configuration for the simulator. This section has been left in this document in case there is an impact on the software architecture.

The following table describes the configuration of switches, jumpers and links on the EL1216 amplifier breakout board.

Table 22: EL1216 Amplifier Breakout Board Configuration

| Jumper, Link or Switch | Position | Function | Description | Comment |
| --- | --- | --- | --- | --- |
| LK6 | 1-2 | Amplifier Remote Hardware Enable input controlled by a sourcing output | Amplifier Remote Hardware Enable input control configuration – floating contacts, sourcing output, sinking output | For 2DFSIM, Remote Hardware Enable input is used by a chassis enable switch supplying 24 Vdc – paralleled between TB2 and TB8 on all boards and front panel switch |
| LK7 | 1-2 |  |  |  |
| JP1 | 2-3 | Amplifier STO input controlled by floating contacts | Amplifier STO input control configuration – floating contacts, sourcing output, sinking output | For 2DFSIM, Jumpered terminal block installed on all boards in TB6 or TB7 |
| JP2 | 2-3 |  |  |  |
| LK8 | 2-3 | Amplifier External Remote Abort input controlled by floating contacts | Amplifier External Remote Abort input control configuration – floating contacts, sourcing output, sinking output | Not used |
| LK9 | 2-3 |  |  |  |
| LK10 | 2-3 | Amplifier External Remote Reset input controlled by floating contacts | Amplifier External Remote Reset input control configuration – floating contacts, sourcing output, sinking output | Not used |
| LK11 | 2-3 |  |  |  |
| SW1 | Up | Local Hardware Enable on | Local Hardware Enable/Disable |  |
| SW2 |  | Local Reset | Amplifier External Local Reset | Momentary |
| SW3 | Down | Local Motion Abort off | Local Motion Abort |  |

# Operational Descriptions

How to use the I/O module Bits

# Reference Documents

| Ref. No. | Document Name | Version/Date |
| --- | --- | --- |
|  | 2DFSIM-MODES OF OPERATION | TBA |
|  | 2DFSIM Equipment IP Addresses | TBA |
|  | http://docushare.aao.gov.au/docushare/dsweb/View/Collection-24120 | Copley Amp Config |
|  | http://docushare.aao.gov.au/docushare/dsweb/View/Collection-19840 | AEROTECH Linear Motors |
|  | http://docushare.aao.gov.au/docushare/dsweb/View/Collection-19839 | X1, X2 & Y encoders |
|  | http://docushare.aao.gov.au/docushare/dsweb/View/Collection-24363 | Jaw Gearmotor |

Unless otherwise noted, TxPDOs (Transmit Process Data Objects) and RxPDOs (Receive Process Data Objects) assume default mapping as per CiA 401.

Unless otherwise noted, TxPDOs (Transmit Process Data Objects) and RxPDOs (Receive Process Data Objects) assume default mapping as per CiA 401.